## Slide 1: The Printed Optical Particle Spectrometer Stratospheric Aerosol Measurements for SAGE III-ISS Validation
Troy Thornberry†, Elizabeth Asher†, Steve Ciciora, Laurel Watts† and Ru-Shan Gao
NOAA Earth System Research Laboratory, Chemical Sciences Division
†and CIRES, University of Colorado-Boulder

## Slide 2: Printed Optical Particle Spectrometer
POPS specifications:
Single-particle detection
140 – 3000 nm diameter range
3 cm3/s sample flow
550 g
5 Watts
Gao et al., AS&T 2016

## Slide 3: Printed Optical Particle Spectrometer
POPS Flow:
Sample flow rate measured using a laminar flow element to provide clean sample flow path
LFE ∆P provides feedback for flow control
Particle free ”sheath” flow set by restriction~ equal to sample flow

## Slide 4: Printed Optical Particle Spectrometer
POPS optics:
405 nm laser diode, ~60 mW
Beam profile shaped by a series of lenses and apertures
Scattered light collected by spherical mirror subtending 38° - 142° (~2.4 sr)
Observed instrument response closely matches Mie theory calculation for spherical particles

## Slide 5: POPS Operation at Stratospheric Pressures
POPS measurement issues at low pressure:
At ambient pressures below ~200 hPa, particle transit through the beam slows, suggesting the inlet jet is diverging
Below ~70 hPa, particle losses were observed as the beam diverged

## Slide 6: POPS Operation at Stratospheric Pressures
Two modifications yielded improved function at low pressure
Install small diameter exhaust nozzle closer to the laser beam (similar to inlet nozzle)
Increased sample flow rate (also increases ”sheath” flow)

## Slide 7: POPS Operation at Stratospheric Pressures
Increased flow and change in geometry both contribute to improved sample jet coherence at low pressure
Optimal performance maintained down to P < 30 hPa, where pump efficiency limits sample flow rate

## Slide 8: POPS Operation at Stratospheric Pressures
Boulder 20190628

## Slide 9: POPS Operation at Stratospheric Pressures
Boulder 20190628

## Slide 10: POPS Particle Size Determination
Calibration with PSLs, n = 1.61 + 0i
Calibration with DMA size-selected dioctyl sebacate, n = 1.45 + 0i

## Slide 11: POPS Particle Size Determination
Mie resonances create non-monotonicity in the Dp vs POPS signal amplitude relationship
Histogram Dp in narrow signal bins and calculate median values
Construct single-valued smoothed curve through median values to estimate Dp from measured signal amplitude

## Slide 12: Index of Refraction
Most OPCs, including POPS, need to assume an aerosol index of refraction as well as spherical shape in order to calculate Dp from scattered light
Leads to potential error in particle sizing

## Slide 13: Index of Refraction
Most OPCs, including POPS, need to assume an aerosol index of refraction as well as spherical shape in order to calculate Dp from scattered light
Leads to potential error in particle sizing

## Slide 14: POPS sonde launches

## Slide 15: POPS sonde launches

## Slide 16: POPS sonde launches

## Slide 17: POPS sonde launches

## Slide 18: POPS sonde launches
UWY 20190828

## Slide 19: POPS sonde launches

## Slide 20: Acknowledgements
Hagen Telg, NOAA ESRL GMD (formerly CSD)
Emrys Hall and Allen Jordan, NOAA ESRL GMD
Lars Kalnajs (CU LASP) and Terry Deshler (U Wyoming)
Wills Dobson, Michael Kotkamp and Richard Querel, NIWA Lauder
Frank Erdesz, NOAA ESRL CSD
NASA SAGE III-ISS Validation Project Marilee Roell, Carrie Roller, Susan Kizer, Travis Knepp, Kevin Leavor